## Slide 1: High dynamic range ion imaging using individual pixel reset of SCAPS detector
ALC’22 Imaging Microscope 19a-3-3
Stigmatic D-SIMS
Naoya Sakamoto (Hokkaido U.)
SCAPS ion detector
24Mg+
108
ions / pixel
Notes: I’m Naoya Sakamoto from Hokkaido University.
We developed a high dynamic range ion imaging system
(10^8を指す)
using a two-dimensional ion detector
(SCAPSを指す)
and a stigmatic SIMS.
(SIMSを指す)
At first, I will explain what kind of imaging we need.

## Slide 2
Sakamoto et al. 2007 Science
Primordial Water
in the Solar System
Cosmic Symplectite（Fe3O4＋FeS）
Heaviest oxygen isotopic composition
Fe/FeS + H2O → Fe3O4
Notes: Fifteen years ago, isotopically anomalous material was accidentally discovered from a meteorite.
These are SEM images and oxygen isotope ratio images.
Scale bars are 5 micron.
This material was found to be oxidized by primordial water in the solar system having a heavy oxygen isotopic composition.
This discovery was entirely coincidental, but open the door to understanding our solar system formation.

## Slide 3
Oxygen
Mass dependent
d17O
d18O
Asteroid
Itokawa
Earth
Yurimoto, Sakamoto et al. 2011 Science
HAYABUSA
???
Required Precision
16O : 99.757
17O : 0.038
18O : 0.205
Isotope abundance
Sun
≒
Bulk Solar System
Sun
Solar Wind(Genesis mission)
McKeegan et al. 2011 Science
Sakamoto&Kawasaki 2019 Metsoc
Yokoyama, Sakamoto et al. 2022 Science
Asteroid
Ryugu
HAYUBUSA2
Solar System Evolution
Primordial Water
Cosmic Symplectite
Sakamoto et al. 2007 Science
Precise isotope imaging
for unknown keystones
Notes: One of the biggest mystery of our solar system is why solar system materials are plotted
(click)
not on the mass-dependent fractionation line,
(click)
but formed a simple strange line.
(click)
Asteroid Itokawa and
(click)
asteroid Ryugu are also plotted on here.
(click)
The primordial water is located on a heavy end member
(click)
and the Sun is another.
(click)
Therefore, we understood this solar system line was formed by mixing of primordial water and the Sun.
In order to trace the solar system evolution occurred in this range
(click)
Precision is required.
(click)
Therefore we need precise isotope imaging to find unknown keystones in somewhere not only high spatial resolution.
My job is to find something beyond human knowledge.
From the next slide, I will explain our strategy for the purpose.
２分３０秒

## Slide 4
©JAIMA
Time of Flight, Ion Trap
Magnetic Sector
m3
m2
m1
Dynamic
Static
1 x 1012
ions/cm2
High
Low
<
<
Primary
Secondary
Pulse
Continuous
1000
1
Duty ratio
1
<
@ frequency 100 kHz
pulse width 10 ns
SIMS
Isotope imaging
Notes: SIMS is a powerful tool for isotope imaging.
(click)
Relative to the static mode, dynamic SIMS continuously acquires signals at a high duty ratio by irradiating a high intense primary beam.
(click)
Therefore, dynamic SIMS is favorable for our purpose.
３０秒

## Slide 5
Probe size
Microbeam < 1 µm
NanoSIMS < 50 nm
FIB, HIM,Cs < 10 nm
Ion optics
High intensity beam
High precision
Wide area
Scanning
Stigmatic
Notes: In terms of imaging method,
The spatial resolution of scanning type is pretty good now, less than 10nm.
But we have to find unknown something that requires precision.
(click)
The stigmatic method has a possibility to make precise isotope analysis for wide area with stigmatic ion optics if we have a good 2-dimentional ion detector.
３０秒

## Slide 6
MCP+FS+CCD system
  Poor Linearity
  Robustness: poor
  Narrow Dynamic Range
MCP+RAE/PSD system
  Good Linearity
  Robustness: good
  Narrow Dynamic Range
  Large dead time
ion
ion
ion
SCAPS system
  Good Linearity
  Robustness: good
  Wide Dynamic Range
  No dead time
SCAPS
Yurimoto et al., 2003 SIA
Notes: Prof. yurimoto and co-workers proposed a 2-dimentional ion imager SCAPS
(click)
having good properties for isotope imaging.
４分

## Slide 7
600x576 pixel
Direct ion detection
Robust
Integral type
No dead time
Non-destructive readout
Nagashima+, 2001; Kunihiro+, 2001; Takayanagi+, 2001; Yurimoto+, 2003;
Sakamoto&Yurimoto, 2006; Yamamoto+, 2010; Sakamoto+, 2012
SCAPS
(Stacked CMOS Active Pixel Sensor)
Notes: SCAPS is a solid state 2 dimensional CMOS type ion detector.
This detector has a pixel electrode in each pixel for direct detection of ions.
This is an output characteristics of a pixel.
You can see this detector has a dynamic range of 4 to 5 order of magnitute for each pixel with good linearity.

## Slide 8
sample
aperture
28Si-
SCAPS
MS
Stigmatic
ion optics
ion source
28Si-
Köhler
Cs+
Dynamic
Stigmatic
SIMS
SCAPS
CAMECA ims series
Notes: Combination of these elements realized precise isotope imaging.
The primary ion beam is irradiated through the aperture to the sample surface by Kohler illumination.
Secondary ions are separated with magnet and transferred by stigmatic ion optics to SCAPS ion imager.

## Slide 9
sample
aperture
28Si-
SCAPS
MS
Stigmatic
ion optics
ion source
28Si-
Köhler
Cs+
Dynamic
Stigmatic
SIMS
SCAPS
4 nA
28Si-
28Si-
40 nA
sub-μm resolution with sub-μA beam intensity
Notes: This is an example of stigmatic isotope imaging.
This is a silicon isotope image with 4nA primary intensity.
The line width of 3 is 3 microns, 2 is 2 microns.
You can see the beam shape is that of the aperture shape.
This is an image taken by 40 nA primary beam.
Although the Kohler illumination is disturbed, we can see that the spatial resolution is almost the same even 10 times higher beam intensity.
(click)
If we enlarge the 40nA image,
You can also recognize 1 micron bars and small scratches.
(click)
This combination allows the use of a strong primary beam while maintaining submicron resolution.
６分

## Slide 10
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)
16O : 99.757 18O : 0.205
Notes: This is an example of high precision isotope imaging for meteorite.
This is O-16 image, this is O-18 and isotope ratio.
Scale bar is 10um and color scale is -60 to 20 permil.
Compared to O-16, O-18 is less abundant.
Therefore the O-18 image is more rough.

## Slide 11
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)
Notes: If we have another images and integrate, the roughness of 18O decreases and you can recognize something by your mind’s eye.

## Slide 12
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)
Notes: We have to accumulate more to see something

## Slide 13
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 14
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 15
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 16
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 17
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 18
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 19
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 20
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 21
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 22
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 23
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 24
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 25
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 26
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 27
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 28
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)

## Slide 29
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
10μm
18O
18O/16O
Analysis
sequence
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
16O
18O
High precision isotope imaging
-60
20
d18O(‰)
Notes: Finally, we can see clear distribution of isotopic composition differences.

## Slide 30
16O
18O
18O/16O
Cumulative
number
1
2
4
8
16
-60
20
d18O(‰)
10μm
10 x 10 pixel
(1 pixel = 0.2 µm)
concentration
Uniform
Cumulative number
10
1
100
1
10
100
Standard Deviation
(‰)
Standard Error
statistical error :
slope 1/2
(‰)
10
1
0.1
OK
Notes: Let’s evaluate quantitatively.
(click)
If we calculate this 10 by 10 pixels, the standard deviation decreases as increasing the cumulative number.
When this 2-miron area is calculated as the standard error, the error reaches less than 1 permil.
(click)
The slope of 0.5 indicates the error is statistically decreasing.
(click),
Therefore, high precision isotope imaging is possible less than 1 permil
(click)
if the concentration is uniformly distributed.

## Slide 31
1H-
Apollo11 rock44
Greenwood, Sakamoto et al. 2011 Nature Geoscience
Water on the Moon
1H-
concentration
% ~ ppm
Dynamic range required for heterogeneous sample
Water on the Moon
1H : 99.9885
2H : 0.0115
Isotope abundance
Notes: How about this case?
This is a moon rock brought back by apollo 11.
As you can see, we found that the lunar apatite crystals contain a small amount of water.
(click)
In the case of lunar water, the hydrogen concentration is very different among crystals.
Moreover, isotopic abundance ratios differ by four orders of magnitude.
(click)
For heterogeneous sample like water, we need high dynamic range for precise isotope imaging shown before.

## Slide 32
High Dynamic Range (HDR) imaging
by SCAPS ion detector
Signal
Time
Pixel Reset
No Reset
( Integration )
Time
High ion count
Low ion count
Destructive readout
(DRO)
Non-Destructive readout
(NDRO)
Notes: This is a concept of high dynamic range imaging by SCAPS ion detector.
Pixels with high ion count are reset periodically, while pixels with low ion count continue to accumulate to avoid reset switching noise.

## Slide 33
Reset
Readout
Vertical Scanner
Horizontal Scanner
Pixel Area
600x576
AOB
Chip Design
Notes: This concept is realized by the pixel structure of SCAPS with separate reset and readout circuits
and individual pixel selection ability based on the XY addressing method.

## Slide 34
Individual pixel reset
Notes: This is a pulse timing chart for one line.
Reset pixels can be selected individually using two reset pulses.

## Slide 35
24Mg+
10μm
0
500
1000
1500
2000
2500
Ion counts / frame / pixel
Pixel
0
100
200
300
400
500
0
1500
3000
Ion counts / frame / pixel
Reference image
DRO
NDRO
thresh
Set thresh
MgAl2O4
CaAl2Si2O8
Ca2(Mg,Al,Si)2O7
Notes: I will explain how to make HDR image using a magnesium isotope map as an example.
Color scale is here.
(click)
White, back and gray colors are corresponding to these minerals.
(click)
This is a intensity graph of this line.
(click)
Set a threshhold value for reset pixel and non-reset pixel.

## Slide 36
0
500
1000
1500
2000
2500
Pixel
0
100
200
300
400
500
DRO
NDRO
Change driving pulses by FPGA
Set thresh
DRO-NDRO map
DRO
NDRO
thresh
Ion counts / frame / pixel
Notes: Here is a reset pixel map.
FPGA code written by LabVIEW software dynamically generates drive pulses.

## Slide 37
DRO-NDRO map
DRO
NDRO
1
10
100
1000
Frame number
10
-4
10
-3
10
-2
10
-1
Output voltage (V /pixel)
10
0
DRO
NDRO
x 1000
Total acquired frame
Notes: If we obtain 1000 image frames, the output signal from the reset pixel is constant and the output signal from the non-reset pixel increases linearly.
(click)
The final integrated signal is used for non-reset pixels, and the signal from reset pixels is multiplied by the number of frames.

## Slide 38
NDRO
1
10
100
1000
Frame number
10
-4
10
-3
10
-2
10
-1
Output voltage (V /pixel)
10
0
10
-4
10
-3
10
-2
10
-1
10
0
10
1
10
2
10
3
10
4
HDR readout (V)
DRO
NDRO
HDR image (Log scale)
x 1000
DRO
Total acquired frame
24Mg+
Stable isotope ratio
Radionuclide dating
Partition coefficient
Dynamics of water by Cryo-DSIMS
108
ions / pixel
Notes: This is a composite image created in that way.
Color is a logarithmic scale from 10 to the minus 4 to 10 to the plus 4.
The signal difference of the squared areas is found to be six orders of magnitude, and even the lower intensity area has a signal of more than three orders of magnitude.
(click)
The high dynamic range ion imaging system allows us to have elemental and stable isotope ratio distributions for highly heterogeneous materials.
The wide range is also effective for precise age dating because the age is calculated by a slope.
This system is helpful for partition coefficient of high pressure materials and the water dynamics in geology and also biology by fully cryogenic dynamic SIMS system.

## Slide 40
クライオ搬送容器
クライオ試料ホルダ
液体窒素デュワー
クライオ同位体顕微鏡
クライオステージ
環境制御型エアロック
Notes: そこで、凍結試料を分析できるダイナミックSIMSであるクライオ同位体顕微鏡を開発しました。
前処理された試料を、クライオ搬送容器で運び、環境制御型エアロックを使って、超高真空の装置内部に導入し、クライオステージ上で分析します。
（試料ホルダを指す）
最も難しいのは、全て凍結させたままで行う必要がある、試料の前処理です。

## Slide 41
クライオ試料前処理システム
クライオワークステーション
クライオ切断研磨装置（開発中）
液体窒素
クライオシャトル
作業スペース
研磨ブレード
自動プランジ凍結装置
液状試料を浸漬凍結
月面での試料調整システムに応用
Notes: そのために、凍結したまま試料を前処理するシステムを着々と準備してまいりました。
生体や有機薄膜デバイス、DAC試料など、様々な試料に対応できるようになっています。
現在、固い岩石を、凍結したまま液体窒素の中で切断および研磨を行う装置を開発中で、研磨ブレードのテストをしている段階です。
（クリック）
液体窒素中では、試料を直接触ることはできませんので、この開発を通して得た経験は、月面での試料調整システムに応用していけるのではないか、と考えています。

## Slide 42
Cryogenic D-SIMS
Freeze
Expose
Au coat
Cryo SIMS Holder
Transfer in Liquid Nitrogen
Cryogenic sample preparation system
Cryo AirLock
Cryo Stage (-190℃)
Analysis (all-night operation available)
Notes: クライオ分析の流れです。
まずクライオ環境下で表面出し、導電膜塗布を行った後、削り出しのクライオホルダーにセットします。
液体窒素で満たした容器でSIMS装置のそばまで運搬し、クライオエアロックに格納してSIMSのクライオステージに導入します。
あとは通常通りに分析することができます。

## Slide 43
Vein
Vein
Vein
100µm
16O-
18O-
18O/16O
Water dynamics in Liver
18O water（H218O）
Highly concentrated in blood vessels
Blood vessels
Blood vessels
Gradual migration through hepatocytes
Notes: こちらは凍結した肝臓の16O、18O、およびそれらの比をとった同位体比のイメージです。
スケールは１００ミクロンです。
こちらの絵で、赤いほど注入した18Oが多く、真っ赤なものは18O水がたくさん通った血管の断面と思われます。
また、こちらの大きめの血管から、中心静脈に向かって流れていき、静脈の周りで濃く集まっていくグラデーションも読み取れます。
このように、クライオ化に成功したことで、水などの揮発成分を、真空でも逃がさずに同位体分析を行うことができるようになりました。
８分（１５分）